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Product Overview
The O2Call Premium is an advanced oxygen monitoring and safety device designed to ensure stable and safe oxygen supply from a high-purity PSA oxygen generator or ambient oxygen source.
It is equipped with a high-precision oxygen sensor and pressure sensor to continuously monitor oxygen purity and supply pressure.
In addition, O2Call Premium detects abnormalities in the power supply of connected equipment (such as the oxygen generator) and provides visual and audible alarms in case of any malfunction.
The system instantly detects critical issues such as decreased oxygen concentration, pressure fluctuation, or oxygen leakage, and triggers alarms to prevent accidents or oxygen supply failure.
For reliable alerting, the device includes a warning light, speaker, and two non-contact voltage alarm outputs for integration with external safety systems.
Key Features:
· Built-in high-precision oxygen and pressure sensors
· Real-time monitoring of oxygen purity, supply pressure, and power status
· Automatic alarm for oxygen concentration drop, pressure anomaly, or power failure
· Equipped with warning light, speaker, and dual non-contact voltage outputs
· Enhances operational safety in oxygen supply environments
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Specifications
	Category
	Specification

	Product Name
	Oxygen Alarm System

	Model
	O2CALL Premium

	Manufacture
	O2MOS

	Country of Origin
	Republic of Korea

	Functions
	Detects and alarms for Oxygen purity drop
Pressure drop
Power outage

	Oxygen Sensor Model
	Zirconia solid electrode

	Oxygen Detection Range (Display)
	0.0% – 99.9%

	Oxygen Detection Resolution
	0.1%

	Effective Detection Range
	75.0% – 95.0%

	Alarm Setting (Purity)
	User-defined between 0.1% – 99.9%
Upper/lower limit alarm output

	Calibration (Purity)
	Adjustable in 0.1% units

	Purity Measurement Cycle
	Every 10 minutes for 10 seconds(Valve open method)

	Detection Range (Pressure)
	0.0 – 3.0 kgf/cm²

	Resolution (Pressure)
	0.1 kgf/cm²

	Accuracy (Pressure)
	±0.1 (≤2.0 kgf/cm²), ±0.3 (≤3.0 kgf/cm²)

	Alarm Setting (Pressure)
	User-defined between 0.1 – 3.0 kgf/cm²

	Power Outage Detection
	Detected within 2 seconds

	Alarm Output
	Red LED blink and built-in speaker alarm
Also includes 2 dry contact signals (AC220V 10A MAX / DC100V 5A MAX)

	Communication
	RS-485

	Oxygen Inlet Port
	Ø6 mm 

	Power Supply
	220VAC, 50/60Hz
Includes 9V rechargeable backup battery

	Power Consumption
	≤ 10W

	Operating Conditions
	0°C – 60°C, 0% – 85% RH

	Dimensions / Weight
	210 × 160 × 300 mm / 3.0 kg

	Material
	ABS enclosure
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