Emergency Oxygen Supply Control Device (Single Type) 
Product Overview
The Vehicle Emergency Oxygen Supply Control Unit is designed to ensure the stable and continuous delivery of oxygen generated by the PSA oxygen generator in mobile environments. It connects directly to the oxygen outlet of the PSA generator and continuously monitors the oxygen pressure level to detect any abnormalities in the oxygen supply.
In the event of an oxygen supply failure or a power issue with the air compressor or PSA generator, the system automatically activates the emergency oxygen supply line, opening the valve of a backup high-pressure oxygen cylinder to maintain uninterrupted oxygen delivery.
Additionally, the unit provides visual and audible alarms to alert the operator of abnormal pressure or power conditions, ensuring safe and reliable operation.
Key Features:
· Continuous monitoring of oxygen pressure and generator power status
· Automatic switchover to emergency oxygen supply during system failure
· Integrated alarm system for oxygen supply and power faults
· Ensures stable and safe oxygen delivery in vehicle-mounted PSA systems
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Product Specifications
	Category
	Specification

	Product Name
	Emergency Oxygen Alarm Device (Single)

	Model Name
	O2MOS-RO2S-SNL

	Manufacture
	O2MOS

	Country of Origin
	Republic of Korea

	Detection Function
	PSA Oxygen pressure detection (6 mm one-touch fitting)
Power abnormality detection (DC power)

	Internal Configuration
	2-port flowmeter (Manual Operation: 1-port, Emergency Operation: 1-port)	
DC power & pressure alarm 
Emergency valve
Includes DC battery for backup power

	Emergency Oxygen Inlet
	10 mm one-touch fitting 

	Emergency Oxygen Outlet
	6 mm one-touch fitting 
Auto supply upon alarm valve activation

	Dimensions / Weight
	340 × 50 × 240 mm / 3 kg

	Material
	STS304



Installation & Wiring Instructions

- Wall-mounted or installed with vibration-damping pad on base plate using direct screws
- Connect pressure sensor and emergency/supplementary oxygen to discharge hose line.
- Connect power sensor wires to compressor power connector.
- Connect emergency oxygen to 2-port flowmeter
Web : www.o2mos.com, www.o2mos.co.kr,  E-mail :  dkahn@dksolutions.co.kr
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